## Slide 1
Designing Highly Efficient LXI-based Systems
Tom Sarfi, VTI Instruments

## Slide 2: The Case for LAN in ATE And making it viable in a T&M environment
The Case for LAN in ATE | 
Comms Type | Serial
Bus Speed | 12.5 MB/s (100 base-T)
 | 125 MB/s (Gigabit)*
Host Interface (Cost) | LAN Port, embedded (NA)
Cabling(Cost) | CAT5e ($5)
Max cable length | 100M-1kM to limitless
Addressing | Static IP - Soft programmable
 | AutoIP
 | DHCP Assigned
Multi-instrument control | Ethernet Switch
Max # of Instruments | NA
Power options | PoE capability
Ubiquity Level | Extremely High/Pervasive
The Need for LXI in ATE | 
Interoperability | Conformance testing required
Software drivers | IVI (per specification)
 | SCPI support (vendor dependent)
Device Synchronization* | IEEE-1588 (PTP)
Instr-Instr Handshaking* | LAN Events
 | Wired-Trigger Bus

## Slide 3: Test Programming Paradigm Communicating between the controller and the ATE
GPIB, cabled PCIe, MXI2, Ethernet, USB
Application Code
C++
LabVIEW
JAVA
Instrument Driver (API)
IVI
VXIplug&play
SCPI command set
Hardware Driver
VISA/Device Driver
DMA
SICL
Transmission Mechanism
GPIB/Word Serial
DMA (PCI-PCI, PCI-VME)
TCP/IP
High degree of variability

## Slide 4: TCP/IP Protocol Overhead Data transmission reliability
Allowable Ethernet Transport Layers in LXI
  TCP – guaranteed data transmission
  UDP – sometimes referred to as “unreliable” protocol, possible but unlikely to lose data packets (LXI specification 3.3.4 Observation)
  More ‘baggage’ required to ensure data delivery using TCP/IP
  IPV6 has larger address fields
IP Header
TCP Header
TCP Data

## Slide 5: LAN Speed Test – Gigabit Link Characterizing the Impact of Overhead
TCP Connection: All measured times are roundtrip, (includes both request and reply).
  Time to open TCP connection is not included.
Client – Server setup
Client sends a 24 byte data packet containing requested data length of reply.
Server returns packet(s) with total data length of 16+requested length.
Frame overhead is 54 bytes (not included in any of the above).
Both NIC’s have TCP Offload capability.
Laptop – Win7
Switch
Switch
Desktop - XP
Total cable length ~ 100 feet
Dual Core 2.53 GHz 4 GB Ram 64 bit OS
Dual Core 3.22 GHz 3 GB Ram 32 bit OS

## Slide 6: Some LAN Speed Test Results – Gigabit Link Carry as much as you can on a trip
Laptop – Win7
Switch
Switch
Desktop - XP
Total cable length ~ 100 feet
Dual Core 2.53 GHz 4 GB Ram 64 bit OS
Dual Core 3.22 GHz 3 GB Ram 32 bit OS

## Slide 7: How an Instrument can Impact Test Time Some measurements take time
Digital Multimeters
  Aperture times (can be 10’s of ms for high accuracy measurements)
  Settling times associated with range and function changes
Mechanical switch settling time
  3 ms – 50 ms depending on application
Simultaneous sampling digitizers
  # of acquired samples/sample rate
Power Supply Settling Times
Send CMD
E.g. Init Measurement
Instrument Action
e.g, arm/trig/integrate
Protocol Overhead
CMD
Query operation complete status
Send CMD
Send CMD
Send CMD
Request Response
e.g; Fetch Data
Protocol Overhead
CMD
Instrument Send Data
Protocol Overhead
Data
In many applications, the overhead is a negligible percentage of overall cycle time
Shipping more data in one packet also reduces the impact on test time
Overly Simplified View of Measurement Cycle Time

## Slide 8: Creating a More Efficient Test System Case Study - Developing More Efficient Instrument Drivers
Replace a MXI2 slot 0 interface with gigabit LXI-VXI interface
  Enables VXI to be used in a distributed environment
  Simplified infrastructure, no PCI plug-in needed
  Drop-in replacement, no code change needed for compatibility
100 kSa/s digitizer in VXI chassis
  16 simultaneous sampling channels, 128 kSa memory depth per channel (4MB total)
Use soft front panel to acquire data and download to PC
  Initial test results were abysmal, 1.28 seconds to acquire the data, minutes to offload
  Looked at underlying code, VISA viIn16 calls were being used to pull data from memory
  The data was being sent to the PC one sample at a time!
  Replaced viIn16 with viMove calls – retrieve data in blocks to reduce protocol hit
  All data returned in less than 1 second, negligible difference between the two interfaces

## Slide 9
Creating a More Efficient Test System Integrate a more powerful infrastructure using extended functions
LXI Extended Functions

## Slide 10
Case Study – Switch/Measure Applications
Temperature/Voltage Monitoring, Continuity/Isolation
Test Case Parameters | 
Host PC | Dell E5520, i5 CPU @ 2.5 GHz
Ethernet Switch | SMC8505T 10/100/1000
6.5 Digit DMM | VTI EX1262
Multiplexer | EX1202/EX1200-3048 (48 ch)
Measurements | 384
DMM Aperture | 1.67 ms
DMM Function | Fixed (DCV)
DMM Range | Fixed (100 V)
Switch Settling | 5 ms (Electromechanical)
48 Analog Signals
6.5 Digit DMM
48 ch. Mux
ENET Switch
Host Controller
INSTRUMENTATION/APPLICATION VARIABLES:
DMM Aperture/Resolution
DMM Function and/or range changes (delays)
DMM Autozero
Switch Settling Time

## Slide 11
Case Study – Switch/Measure
Software/Application code pacing
Send Command Close ChN
Wait for Switch Settle
Send DMM Measure Command
Poll on DMM Measure Complete
(or wait)
5 ms
1.67 ms
Variable
Variable
Actual Measured:
Protocol overhead carried for every switch measure loop
Removing all variables, best case scenario
1 / (1.67 ms + 5 ms) * 384 = 2.56 seconds
150 readings/second
Send Command Open ChN
Variable

## Slide 12: Creating a More Efficient Test System Integrating Intelligent Instruments
Integrated Switch/Measure into one instrument
  Modular DMM and Multiplexer
  Scanning sequence can be built using graphical interface, or through API
  Scan list resides in on-box memory
  On-box processor paces the sequencing
  The only host intervention that is required is to initiate the scan and retrieve the data
  Internal analog bus reduces external wiring
Actual Measured: 150 rdgs/s
There is minimal impact due to protocol overhead

## Slide 13
The LXI Wired Trigger Bus (WTB) Extension
Emulating modular architecture backplanes
Analogous to VXI and PXI trigger bus
  Familiar hardware-based model for triggering and handshaking between discrete modules
  Eight (8) mLVDS lines with wired-or operation
  Supports segmented trigger bus lanes
    Systems can extend beyond the eight trigger lines
Separate scan list in DMM and Switch
  Sequencing paced by trigger pulses passed across WTB when measurement is complete or switch has settled
  The only host intervention that is required is to initiate the scan and retrieve the data
  LXI0 = MeasCompl -> advances to next element in switch scan list
  LXI1 = SwitchSettle -> triggers DMM to take measurement
WTB Termination
Actual Measured: 145 rdgs/s
Again, there is minimal impact due to protocol overhead

## Slide 14
LXIsync LAN Event Triggering
Peer-Peer messages over CAT5e
LXI LAN Events
  Peer-peer multicast architecture
  LAN messages are sent over the wire as triggers
  DMM sends “Measure Complete” event LAN0 which the switch uses to advance scan list
  Switch sends “Switch Settled” event LAN1 which the DMM uses to initiate a measurement
  LAN signals are converted to backplane signals within each instrument
IEEE-1588 implementation
  Events are sent with timestamps, each instrument uses the timestamp as the trigger event
  Synchronization typically around 10s of us
  Initial results were TERRIBLE! (Lesson learned – l started the test right after powering up the box – did not give the two instruments a chance to become synchronized)
Actual Measured: 92 rdgs/s
Processing LAN events/implementation impacting efficiency

## Slide 15
Comparing Results
Choose the best fit for the application
Method (Electromechanical, 5 ms settle) | Measured
Software Pacing | 63 rdgs/sec
Integrated/Internal Scan Engine | 150 rdgs/sec
LXI WTB Handshaking | 145 rdgs/sec
LXI LAN Event Handshaking | 92 rdgs/sec
Theoretical Max | 150 rdgs/s
Method (Solid-State, 0.2 ms ON time) | Measured
Software Pacing | 60 rdgs/sec
Integrated/Internal Scan Engine | 535 rdgs/sec
LXI WTB Handshaking | 522 rdgs/sec
LXI LAN Event Handshaking | 123 rdgs/sec
Theoretical Max | 535 readings/s
Determinism | 
Software Pacing | Below Average
Integrated/Internal Scan Engine | Excellent
LXI WTB Handshaking | Excellent
LXI LAN Event Handshaking | Average
Resolution of WAIT statement

## Slide 16
Summary
TCP/IP protocol carries overhead with every message sent over the wire
  Testing shows about 300 us per message
  Using UDP can reduce this by an order of magnitude, but data delivery not guaranteed
For many applications, this can be imperceptible
Smart instruments minimize host intervention
Instrument driver structure/application code can have a significant impact
  Incorporate the most efficient programming model available
  LXI extended functions can be used to overcome the impact of overhead
    LXI WTB offers backplane-like performance
    LXI LAN Events can be efficient for distributed applications
    LXI extended functions can be used to overcome the impact of overhead
    LXI-based test systems can be architected for efficiency